## Slide 1: How to do NPGS (Nabity* Pattern Generation System) Electron beam lithography
Pete Eschbach
OSU EM Facility
Joe Nabity, the creator of the NPGS, was here August 28 – Sept 1st 2016. He upgraded the software, tuned up the beam blanker, and updated system settings. We then wrote 20 to 25 nm lines in PMMA with the Quanta 3D and so we can assure you the system works very well using Joe’s setup and the run files of those dates. Align.rf6 also works very well. Pete has tested it since that time.

## Slide 2: Insert Lithography holder with scratched PMMA and focus area and Black hole (Faraday cup).
You will need these tools: screwdriver and our high quality diamond scribe (Teresa’s office)

## Slide 3: Find scratch tip at low current and low mag, write down the coordinates

## Slide 4: Use XT align to rotate stage to make the cleaved edge Horizontal – aids in finding your patterned area(s) later
Cleave face
Al post
PMMA on Si

## Slide 5: Start NPGS, hit 2 on KVM (keyboard, video mouse) switch. Click NPGS on monitor of second computer
It is normal to see this black window, it means the Digital to Analog Converter (that converts your pattern to a writable pattern) is calibrating.

## Slide 6: Focus on gold TEM grid at 10 mm working distance and link

## Slide 7: Now type in 5.5 mm on Z height, watch camera . Now unlink stage!

## Slide 8: In NPGS software click on 6 mm WD button
This sets the microscope working distance to 6 mm

## Slide 9: Use run file editor to find your pre set current for PMMA, generally 35 pA. Select this aperture and 30kV. Thirty kV gives a very narrow neck of scattering in 100 nm thick PMMA.

## Slide 10: Now focus with Z height on CCD camera page using mouse or alternately using Z knob on stage. Update stage position for Gold.

## Slide 11: Now drive to Gold TEM grid for fine focus and Stigmate carefully at 100,000x or above. Optional use Pg_Optimize at 400kx or above to arrive at the best stigmator values.
This was 200,000x image. The gold particles should be very sharp, even the smallest ones that are 10 nm in diameter.

## Slide 12: Mark the gold stage position. You should not change the stigmation values from now on.

## Slide 13: Drive stage to Farday cup, think of where it was relative to gold when loaded then rotate that position in your mind 90 degrees

## Slide 14: Mag up on Faraday cup and measure current. You should first hit Auto Zero on Keithley ebeam electrometer with beam not on cup
Almost enough magnification
This is beautiful, you are in the center of Faraday cup, a black hole

## Slide 15: Enter the measured current value in Run file editor for your file (click on pattern) in NPGS. Should be close to 35 pA for most fine PMMA work.

## Slide 16: NPGS can signal average the Keithley values

## Slide 17: Set current value to what was measured. You can set all patterns at once by entering in one pattern, then double click. It should prompt you to enter for all.

## Slide 18: Magnification back down to less than 50

## Slide 19: With low mag 35x for example, drive to your crack tip(use your x,y value from page 3). Don’t use a preset that would move your stage back to 10 mm – think before using any preset. Quickly mag up and check focus but not stigmation on scratch tip.

## Slide 20: Check Focus at tip at around 10,000x

## Slide 21: Move away (down or up) one frame at 10kx to get away from scratch debris to a clean field. Pause beam

## Slide 22: Before burning contamination spot turn EIS, scan interface to External Escan on. Beam blanker off!
Beam blanker on external and on

## Slide 23: Contamination spot, click on 30 sec spot in NPGS and put green arrow in middle of screen
This is a suitable 30 sec contamination spot
After counting 30 sec, NPGS will mag up for you and turn on beam blanker so you can see the spot. Focus if needed.
Click here!

## Slide 24: Turn off EIS Focus on contamination spot

## Slide 25
Make a 1 sec contamination spot if you are making fine lines.
This is a very good 1 second spot and these beam conditions will result in very good fine lines.
The 30 second spot

## Slide 26: Leave beam running on contamination spot with a scan rate greater than or equal to 300 nSec and now turn on ESI to Ext Escan and turn on beam blanker as in next two slides

## Slide 27: Check beam blanker on external and purple light is lit.
Purple is good.

## Slide 28: Beam blanker, set at 4. Default value, 3, 2, 1 are more narrow settings for ultra fast writing. We have this blanker adjusted at 4, please leave at 4.

## Slide 29: Check beam External Scan Interface

## Slide 30: Ebeam should now be blanked and you can simulate your run file and make last minute changes to run file before moving to starting position and running your file. It is getting close now!

## Slide 31: Now we move to fresh area, suggestion: move down two at 1000x because scratch tip is usually up. Then move one frame left because Nabity patterns generally move Counter clockwise.
Scratch tip
Move two fields down one right before starting write

## Slide 32: After moving to fresh area, record microscope positions

## Slide 33: Now select .rf6 file and click process run file in Nabity software.

## Slide 34: You will see the microscope image on the PC display flicker but it is not in synch, this is normal
Black screen that doesn’t change is not normal

## Slide 35: Follow prompts, space bar to continue, n if you don’t want to repeat pattern.

## Slide 36: Hit any key to keep going

## Slide 37: It is now writing. With direct memory access, NPGS will write the pattern to the screen faster than the beam will move. This is normal.

## Slide 38: Next pattern is written, you should here the stage move and a bleep between patterns.

## Slide 39: Final entity of a run file is comment section asking you a few questions.

## Slide 40: When done, record stage positions, they should match the beginning. If not, let us know!
Also, close NPGS if you are done.
Why close it?
Doing this will prompt NPGS user to select the correct project, directory and will also run the DAC calibration.
Also switch KVM (keyboard, video, mouse) switch to position 1.
Turn beam blanker off.
Return ESI box to all out positions (so no LED’s are lit).

## Slide 41: You can see a latent image by using a very low current and quick scan

## Slide 42: Develop your PMMA we have developer, stop bath and remover here in great quantity!
Joe Nabity recommends 60 to 70 seconds in developer (MIBK/IPA). 60 for 22C room. 70 seconds if room is cold, say 18C.
Quickly move to stop bath,30 seconds Isopropyl Alcohol.
Rinse in DI water for a minute and blow dry.

## Slide 43: After you have confirmed that patterns are on your sample you may wish to attempt lift off or metallization of areas where pattern is not.
In lift off the sample is coated at a small angle with a point source metal such as gold or Chromium in the EM Lab PVD system (old Varian vacuum evaporator).
Acetone (20 minutes) or NMP (Pyridinone, 30 seconds) is then used to strip off the PMMA areas and what is left behind is a metallized area where PMMA was not. You can then align subsequent patterns to these patterned metal lines and build devices by repeating the photo, dep etch process.

## Slide 44: Wagon Wheel after exposure to PMMA with NPGS, then Chromium coating and strip with Acetone to reveal fine Chromium lines.

## Slide 45: Align flow process from Joe Nabity
Run DevPat1 with NPGS
Develop Sample
Evaporate metal on Sample
Do Lift-Off
Run** Device2withAlignment with NPGS
Spin resist on Sample
Spin resist on Sample
Develop Sample
Evaporate metal on Sample
Do Lift-Off
Blue steps done in SEM
**”Align” and “DevPat2” are combined in this step.
1st Processing Run
2nd Processing Run

## Slide 46: After developing away exposed areas from DevPat1.RF6 lithography run.

## Slide 47: Now we will use NPGS to align to these marks

## Slide 48: After exposing resist with DevPat1 and your device in middle of this align pattern. Align to the alignment marks from DevPat1.
A. Do your best alignment with stage and image from SEM. B. Then Run Align.rf6 and image of your pattern will be acquired.
C. Now move cursors NPGS puts up with edge of SEM pattern that NPGS has imported for you!
D. Hit R on screen, calculate rotation matrix
The image area is cleverly only the small squares and so most of resist is not exposed!

## Slide 49: Repeat this procedure with all four alignment marks.

## Slide 50: Final alignment, on these rather poor marks but it worked. You are now ready to run DevPat2 with your pattern in center and all will be aligned with respect to first pattern.

## Slide 51: How to spin coat PMMA onto Si
Recipe 10 at ONAMI/MBI works well. 10 sec spin at 200 rpm followed by a 1:10 sec slow acceleration to 5000 rpm (I stop at 4000 rpm or about 1 minute in).

## Slide 52: Add plenty of PMMA, otherwise you will get a streaky coating and regions that are not coated.

## Slide 53: After wafer gets to 4000 rpm the PMMA 950 should give wafer a blue green tint.

## Slide 54: Bake the wafer at 150C to 180C for 2 minutes. Remove with wafer tweezer as they are hot!

## Slide 55: Storage of PMMA baked wafers
Joe Nabity says spun and baked PMMA wafers will last for years if stored in lightproof dessicator.
We believe this as we stored some spun wafers in a non light proof not dessicated container for one year and they still worked although they were degraded.
So, light proof and dessicated is the best recommendation.
You can always test your PMMA by putting a test strip in developer for 60 seconds then stop in IPA The unexposed test strip should remain blue to green color after this development and stop.

## Slide 56: Writing from BitMaps you find on internet.
Find a bitmap on internet. Right click it, save it
Re open it in paint and save as .bmp 256 colors as NPGS will not work for .jpeg or .tif

## Slide 57: Moving it from Internet Computer to the support PC where Nabity resides
Support PC is not connected to internet directly. But it shares a common mapped drive, Archive, that connects the report (third computer that is on internet).
Move file from report PC to Archive drive on Support PC.
Then open Design cad on Support PC and import from file: import
See next page

## Slide 58: Importing file into Design CAD

## Slide 59: Now center and save from NPGS menu in Design Cad
Size your import, drag cursor to corners of where you want image
Hit MM icon in the Design Cad launched from NPGS. This will center your pattern.
Save the pattern as .dc2 from NPGS menu not the design cad File menu!
Next page for visual

## Slide 60: How to import
NPGS menu to save!
File menu to import
Mark where you want it.
Hit MM to center

## Slide 61: Converting to RF6
Right click on DC2 file
Or, grab an old RF6 with a similar pattern. Open it up with run file editor click on pattern and insert new file name of your bitmap

## Slide 62: Open existing RF6, double click on pattern and select your new freshly created .dc2 of your bitmap

## Slide 63: Now set your doses for ion beam: 1 mC/cm^2 is ½ micron in Silicon.

## Slide 64: Dose scheme min (dark) .1 mC/cm^2 max 2 mc/cm^2 (deeper should be lighter colors)

## Slide 65: ESI in these positions for ion beam lithography

## Slide 66: Beam Blanker off for ion beam lithography

## Slide 67: Simulate pattern

## Slide 68: Then run!
You will most likely have to adjust the size and depths of your image.
Remember to focus ion beam well
NPGS does not control ion beam automatically, so you will need to enter the closest mag to correct before!
Electron lithography is remotely controlled for mag and current with “auto it.”
Cut ion image at 52 degree
Then electron image your lithograph at 0 degrees.

## Slide 69: Result
Ion beam lithograph
Original