## Slide 1: Using The Struers Tenupol 3 Twin Jet polisher in the Santala Lab.

## Slide 2: Twin Jet Polisher* in the Santala lab. *A twin jet polisher is probably the fastest way to prepare metallurgical samples. You simply need to cut a thin slice on diamond saw. Polish to .5 mm thickness and punch out a 3 mm disk . Polish jigs and 3 mm punch are available in the OSU EM facility if needed. Pete Eschbach can assist you with mounting and cutting on diamond saw, polishing to .5 mm thickness and punching your 3 mm washer. And Pete can help with the twin jet polisher as it is a joy to behold!
Struers Tenupol 3. on semi permanent loan from DOE NETL

## Slide 3: Add your etchant. This is recipe A2: 70% ethanol, 10% perchloric acid, 10% ButoxylEthanol, and 10% DI water. Use proper PPE and splash resistant glasses!

## Slide 4: Get yourself a good work area and some filter paper to collect samples.

## Slide 5: Load your 3 mm diameter foil

## Slide 6: Set the parameters for milling. A good starting point at room temperature is: Photosensitivity 8, Voltage 35, current display .5 nA, polish of 2 minutes and flow rate 3
More photo sensitivity corresponds to thinner samples.
The setting here, 4, is too aggressive at room temperature and will obliterate your foil! Try 8 to start off and work down to 6 if need be.

## Slide 7: Now push the mill button. It will turn red and you will hear the jet pumps start. Eventually you will see the etch current, nominally .4 amperes.

## Slide 8: When the IR photo sensor detects that it is done, a hole has been generated, the twin jet stops and the light turns green.

## Slide 9: Save your solution in a marked container and with secondary containment.
Marked container!
Put in secondary containment if you are transporting it.

## Slide 10: When done run water through the system. This will require that you put a dummy foil in so that conductivity is sensed and jets turn on.