## Slide 1: How to make a low cost* TEM wedge sample with grinder polisher
Peter Eschbach
Oregon State University
* This is a much lower cost alternative to a Million Dollar FIB prep. It is a technology you can take with you anywhere in the world. You only need a grinder polisher $4000, $1000 diamond saw, and a $1500 tripod polisher and access to a good low power stereo microscope. And roughly $200 worth of diamond lapping films.

## Slide 2: Protect glass cover slip is applied to your die or metal foil with 110 adhesive
sample
Glass cover slip with 110 adhesive, Mbond works well also

## Slide 3: Put on hot plate at prescribed temperature and squeeze sample to cover glass with “alligator clip.”

## Slide 4: Protect glass is now bonded as indicated by amber color of adhesive.

## Slide 5: There are two options for ground samples, cross section and plan view. This is plan view and it has no cover slip, no protect glass. First we will do Cross section
Cross sectioning paddle
Sample to be thinned plan view

## Slide 6: Sample prep method 1, cross sectional
Cross sectional is a bit more complicated
We first mount sample parallel to long axis of cross sectioning paddle and then grind and polish to the area of interest (AOI).
Then we remount the sample perpendicular with AOI down.
With AOI down we thin to a wedge for electron transparency

## Slide 7: Cross sectional sample prep Overview
Step 1 protect layer
Glued on
sample
Area of interest. For example a defect in metal
Protect glass
Thin adhesive bond line(s)
Grind and polish to area of interest (AOI)
Step 2: mount long axis // to cross section paddle and grind to area of interest
Step 3 sample de-bonded and rotated to thin to AOI
Less than 3 mm
For sensitive samples a second protect glass is added to bottom
Optional second protect glass
AOI down!
Towards glue
2
3
In step 3 we have unglued and re-bonded to thin a wedge to area of interest
2
3
AOI

## Slide 8: Work must be done on diamond saw to trim piece
Trim your piece on high speed diamond saw

## Slide 9: Here we are slicing a thin slice of Nickel steel so we can then planarize it on grinder.

## Slide 10: Planarizing Nickel Steel sample
Sample
Fischione model 160 planarizing jig
This area captures sample for holding it!

## Slide 11: Now, we can hold onto the sample and decrease thickness with knob as we grind!

## Slide 12: Now that sample is planarized, metal samples typically don’t come planar, Silicon does
We must trim it on diamond saw to maximum dimension 3 mm so that we can glue it to a 3 mm TEM grid. No standard TEM anywhere in world takes a sample bigger than 3 mm.
Then we mount it with hot melt, actually tree resin
Then we can thin it to rough electron transparency with the tripod polisher tool and diamond lap films (DLF) in LPSC room 145.
Finally, PIPS, precision ion polishing (PIPS) in Melissa Santala’s lab in Dearborn.

## Slide 13: Our planzarized Nickel metal foil now being trimmed to 3 mm max dimension

## Slide 14: Here we are trimming an integrated circuit so that maximum dimensions are 3 mm, so it can be thinned and glued to a 3 mm TEM grid.

## Slide 15: Level the tripod polisher and grind the Delrin feet to planarize. Per the instructions in the Allied High Tech Tripod polisher in LPSC room 145.

## Slide 16: Slight angle on tripod polisher so that a wedge is made
At start
At end
We thin through glass and leave just region of interest!
Start with 30 micron DLF
Then 15 um
6 um, 3um, 1/2 um
And final polish at .05 micron Silica or Alumina

## Slide 17: Once we have thinned it to rough electron transparency, 150-300 nm we mount it on a grid. This is the hardest part.
Float the sample from the hot melt resin by soaking in acetone in a small beaker with filter paper at bottom
Collect sample in filter paper
Then put small drop of 110 or M bond adhesive on TEM grid of your choice, author believes Slotted grid with tab is easiest, see next slide.
Drop grid wetted with drop of epoxy onto sample on filter paper, do not attempt to pick up sample up with tweezer. It is very fragile!

## Slide 18: Sample pick up onto slotted grid
Glue slotted grid to glass slide with Loctite removable adhesive or hot wax.
Put dab of 110 adhesive on slotted grid.
Lower slotted grid with glue side down onto the thinned sample.
Put slotted grid and glass slide on hot plate to cure 110 adhesive
Cut tab off slotted grid to and place grid in TEM grid box for analysis!

## Slide 19: Mounting the wedge sample to TEM grid, higher Mag image
This Si sample was still attached to the glass protect layer. To make a viable sample, it would have to have been separated from the glass backing then glued to the slotted grid.

## Slide 20: Plan view sectioning, method 2
Sample is mounted just to the edge

## Slide 21: Switch to 15 um Diamond lap film

## Slide 22: Thinned about half way.

## Slide 23: Then switch to 3 um Diamond Lap Film
Silicon sample is quite thin and barely visible

## Slide 24: Now switch to ½ um DLF and then final polish with .05 micron Silica
Amber color of Si indicates electron transparent, and you can see electronic interconnects. Sample is ready to be debonded and transferred to a TEM grid

## Slide 25: Once electron transparent wedge has been formed, time to transfer it to a TEM grid
Follow same procedure detailed in slides 17, 18
Keep in mind that the thin side of sample should be in middle of grid when you glue it for maximum viewing potential at tilt.